Table 2S. Summary table of the test procedures carried out. Numbers denoted custom chambers vs.: (1) LI-8100; (2) EGM-4; 
(3) LI-6400 conifer chamber.

| Test | 668 cm3 custom chamber | 2,385 cm3custom chamber |
| --- | --- | --- |
| Preliminary best configuration options |  |  |
| Test 1: airflow | 700–800 μmol s–1 (2) | 700–800 μmol s–1 (1) |
| Test 1: additional fan speed | Not recommended (2) | 0.7 m s–1 (1) |
| Test 1: internal fan speed | LI-6400 high speed option (2) | LI-6400 high speed option (1) |
| Test 2: system mode | Open system (1) | Open system (1) |
| Improved comparison tests |  |  |
| Test 3: net photosynthesis measurements | Not implemented | R2 = 959; slight overestimation (3) |
| Test 4: respiration measurements | R2 = 0.929; 1:1 relationship (2) | R2 = 0.962; slight underestimation (1) |
| Testing the influence of the chamber on micrometeorological variables |  |  |
| Test 5: pressure changes | Not implemented | Not relevant with the selected settings |
| Test 6: temperature (T) and RH changes | T no implemented; RH rises in some circumstances | Maximum increase of 5°C; not relevant RH increase |